## Slide 1
Please note that for new colleagues attending this meeting, we will be adding your contact details, as provided, to our internal SWRI address database, https://www.swri.co.uk/privacy-policy.html
The webinar will be recorded and placed on the SWRI Extranet for future viewing.
The webinar will be held in accordance with the Competition Law requirements of normal SWRI committee meetings. Members must not discuss any commercially sensitive information, including but not limited to information which is confidential to their own company and/or is relevant to the commercial strategy or activities of their company. This restriction applies during the formal meeting and in any discussions or exchanges around or relating to the meeting. Members are in the best position to judge what is commercially sensitive or confidential and so the primary responsibility for compliance lies with members. Failure to comply may have serious consequences for the individuals involved, their companies and the SWRI itself. Such consequences can include substantial fines and the imposition of criminal penalties including imprisonment.

## Slide 2: Quality Control
Hannah Ford
SWRI Service Group

## Slide 3: Summary
PART 1:

## Slide 4: General Lab Checks Balances
Internal vs external

## Slide 5: General Lab Checks Pipettes
Internal vs external

## Slide 6: General Lab Checks Volumetric Flasks

## Slide 7: General Lab Checks Thermometers
Notes: Internal vs external

## Slide 8: General Lab Checks Other Temperature Loggers

## Slide 9: General Lab Checks Instruments – External/Preventative maintenance
Service contracts

## Slide 10: General Lab Checks Instruments – Gas filters

## Slide 11: General Lab Checks Chemicals and Standards
Notes: CoA
Expiry dates
Quality/purity (mobile phases, solvents)
Extension when necessary – how to justify. What’s appropriate to each lab (vs eg cost of chemical)

## Slide 12: calibrations
Internal and external standards
Notes: Checks in place for new ones
Deuterated IS

## Slide 13: Analytical Sequences – Inclusion of QC Samples
Why we use QC samples
What should the QC sample be
  Storage/expiry/stability
  Multiple/different matrices/levels
  Compromise on ideal QC
Natural levels vs spiked recovery (SQC vs AQC)
Where to put QC samples in your sequence
Bracketing standards
Non-routine methods – spike
  Independence of spiking standards

## Slide 14: Walter A. Shewhart 1891 - 1967
www.magazine.amstat.org
Quality control in manufacturing
Devised the control chart system in 1924
“Assignable-cause” vs “chance-cause” or “Special-cause” vs “common-cause”
A process in “statistical control” has ONLY common-cause variation
Notes: In 1924 Shewhart was working for the Inspection Engineering Department of Western Electric, which manufactured hardware for the Bell Telephone Company. Quality control at the time was limited to inspecting finished products and removing defective items. The company had realised that adapting processes reactively to every non-conforming product was actually counter-productive – increasing variability and resulting in more defections. So Shewhart set out to devise a system to distinguish between what he called “assignable-cause” and “chance-cause” variation (these terms were later on renamed to “special-cause” and “common-cause”) – this system being what we would recognise today as a control chart.

## Slide 15: The First Control Chart 1924
www.deming.org
A few days ago, you mentioned some of the problems connected with the development of an acceptable form of inspection report which might be modified from time to time, in order to give at a glance the greatest amount of accurate information.
The attached form of report is designed to indicate whether or not the observed in the percent of defective apparatus of a given type are significant; that is, to indicate whether or not the product is satisfactory. The theory underlying method of determining the significance of the variations is the value of p is somewhat involved when considered in such a form as to cover practically all types of problems. I have already started the preparation of a series of memoranda covering these points in detail. Should it be found desirable, however, to make use of this form of chart in any of the studies now being conducted within the Inspection Department, it will be possible to indicate the method to be followed in the particular examples.
– – W. A. Shewhart
Notes: Other applications – process quality control also applicable.
How much does Barry know about process control charts (as opposed to analytical) and software packages used? Richard Cowley @ Diageo (Leven pilot plant) Jane Millar

## Slide 16: Control Chart Setup
Notes: How much data is needed for meaningful
Running in new sample alongside old
When to restart a chart

## Slide 17: Control Chart Setup
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
+3SD | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728
Result
Repetition number (or date)

## Slide 18: Some Stats Recap
Mean Average (μ) or Arithmetic Mean
                μ =
  xi is each result, N is total number of results
Standard Deviation (σ): measure of amount of variation in a set of results

## Slide 19: Normal Distribution
Chart: 
Chart: 
| 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | 41 | 42 | 43 | 44 | 45 | 46 | 47 | 48 | 49 | 50 | 51 | 52 | 53 | 54 | 55 | 56 | 57 | 58 | 59 | 60 | 61 | 62 | 63 | 64 | 65 | 66 | 67 | 68 | 69 | 70 | 71 | 72 | 73 | 74 | 75 | 76 | 77 | 78 | 79 | 80 | 81 | 82 | 83 | 84 | 85 | 86 | 87 | 88 | 89 | 90 | 91 | 92 | 93 | 94 | 95 | 96 | 97 | 98 | 99 | 100
SD = 10 | 1.4867e-07 | 6.3698e-07 | 2.4942e-06 | 8.9262e-06 | 2.9195e-05 | 8.7268e-05 | 0.00023841 | 0.00059525 | 0.0013583 | 0.0028327 | 0.0053991 | 0.0094049 | 0.014973 | 0.021785 | 0.028969 | 0.035207 | 0.039104 | 0.039695 | 0.036827 | 0.031225 | 0.024197 | 0.017137 | 0.011092 | 0.0065616 | 0.0035475 | 0.0017528 | 0.00079155 | 0.00032668 | 0.00012322 | 4.2478e-05 | 1.3383e-05 | 3.8535e-06 | 1.0141e-06 | 2.439e-07
Result or variable
Frequency of occurrence
Notes: If I were to poll all of you for your shoe size, I would probably create a histogram that looks a little like this, and this is what we call a “normally distributed” variable. There will be more of you with a close to average shoe size and far fewer with very small or very big shoes. Shoe size is a discrete variable, as each shoe size jumps by at least half a size to the next one, so it only approximates a smooth curve.

## Slide 20: Normal Distribution
Chart: 
Also known as Gaussian distribution or a “bell curve”
Symmetrical distribution with exactly half of results on each side of mean (no bias)
Mean = Median = Mode
Notes: If I were to create a similar plot of all of your heights however, as a continuous variable it would give a smooth curve, and the more data points you have the more smooth the curve becomes. The exact shape of this distribution is called Normal Distribution – there are other shapes of distribution and not every dataset will fit a normal distribution, but it’s this one relevant to us because measurement errors do fit this curve.

## Slide 21: Normal Distribution
Chart: 
Smaller standard deviation: taller and sharper curve
Notes: A smaller standard deviation will result in a narrower distribution where more results are clustered closer to the mean, giving a taller and sharper curve.

## Slide 22: Normal Distribution
Chart: 
Smaller standard deviation: taller and sharper curve
Larger standard deviation: shorter and broader curve
Notes: …and a larger standard deviation conversely will give a shorter and broader curve. This means the likelihood of results falling quite far from the mean increases.

## Slide 23: Normal Distribution
Chart: 
-3SD | 0 | 0.06
-2SD | 0 | 0.06
-1SD | 0 | 0.06
+1SD | 0 | 0.06
+2SD | 0 | 0.06
+3SD | 0 | 0.06
68.27%
± 1 standard deviation from the mean: 68.27% of all results
ie. for every ~3 results, expect 1 to be outside 1SD
μ - 1σ
μ + 1σ
μ
Notes: As the shape of the bell curve is a function of the standard deviation, the values of the mean plus or minus standard deviation will always fall at the same relative point on the curve. This middle section of the bell curve, between the points that are minus and plus one standard deviation from the mean (shown by the green lines), encompasses 68.27% of all the results. So roughly one in every three results will fall outside these green lines.

## Slide 24: Normal Distribution
95.45%
± 2 standard deviation from the mean: 95.45% of all results
ie. for every ~22 results, expect 1 to be outside 2SD
Chart: 
-3SD | 0 | 0.06
-2SD | 0 | 0.06
-1SD | 0 | 0.06
+1SD | 0 | 0.06
+2SD | 0 | 0.06
+3SD | 0 | 0.06
μ - 2σ
μ + 2σ
μ
Notes: …but note that 1 in 22 refers to just results that are a little further from the mean purely due to random variance/fluctuations in the results. So be aware there might be other results that are outwith these points due to actual error.

## Slide 25: Normal Distribution
99.73%
± 3 standard deviation from the mean: 99.73% of all results
ie. for every ~370 results, expect 1 to be outside 3SD
Chart: 
-3SD | 0 | 0.06
-2SD | 0 | 0.06
-1SD | 0 | 0.06
+1SD | 0 | 0.06
+2SD | 0 | 0.06
+3SD | 0 | 0.06
μ - 3σ
μ + 3σ
μ
Notes: You will appreciate that this is a much rarer occurrence – unless you have many many data points, a result that falls outwith these points is very likely to be erroneous and NOT a result of random chance.

## Slide 26: Control Chart Setup Normal Distribution Principles in a Control Chart
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85
-3SD | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272
-2SD | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348
+2SD | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652
+3SD | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728
Result
Repetition number (or date)
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ

## Slide 27: Control Chart Setup Fixed Mean vs Rolling Mean
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 116 | 99 | 102 | 96.5 | 100.2 | 102.8333 | 102.5714 | 103.5 | 101 | 100.3 | 100.1818 | 101.3333 | 101.8462 | 101.6429 | 100.6 | 101.4375 | 100.5294 | 99.9444 | 101 | 101.5 | 101.5714 | 101.3182 | 100.6522 | 100.0417 | 100.6 | 101.2692 | 101.3704 | 101.9286 | 101.5172 | 101.9667 | 102.1935 | 102.4375 | 102.5455 | 101.8824 | 102.3429 | 102.4167 | 102.2432 | 101.7895 | 101.9487 | 101.85
-3SD | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272
-2SD | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348
+2SD | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652
+3SD | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728

## Slide 28: Control Chart Setup Fixed Mean vs Rolling Mean
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 116 | 99 | 102 | 96.5 | 100.2 | 102.8333 | 102.5714 | 103.5 | 101 | 100.3 | 100.1818 | 101.3333 | 101.8462 | 101.6429 | 100.6 | 101.4375 | 100.5294 | 99.9444 | 101 | 101.5 | 101.5714 | 101.3182 | 100.6522 | 100.0417 | 100.6 | 101.2692 | 101.3704 | 101.9286 | 101.5172 | 101.9667 | 102.1935 | 102.4375 | 102.5455 | 101.8824 | 102.3429 | 102.4167 | 102.2432 | 101.7895 | 101.9487 | 101.85
-3SD | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272 | 64.0272
-2SD | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348 | 76.6348
+2SD | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652 | 127.0652
+3SD | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728 | 139.6728
mean at end | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85 | 101.85

## Slide 29: Control Chart Setup Fixed Mean vs Rolling Mean
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 116 | 99 | 102 | 96.5 | 100.2 | 102.8333 | 102.5714 | 103.5 | 101 | 100.3 | 100.1818 | 101.3333 | 101.8462 | 101.6429 | 100.6 | 101.4375 | 100.5294 | 99.9444 | 101 | 101.5 | 101.5714 | 101.3182 | 100.6522 | 100.0417 | 100.6 | 101.2692 | 101.3704 | 101.9286 | 101.5172 | 101.9667 | 102.1935 | 102.4375 | 102.5455 | 101.8824 | 102.3429 | 102.4167 | 102.2432 | 101.7895 | 101.9487 | 101.85
-3SD | 26.8751 | 48.6708 | 41.8648 | 46.7697 | 51.2746 | 55.4591 | 59.1765 | 53.8274 | 55.3322 | 57.5054 | 58.9198 | 60.8612 | 62.1997 | 60.7069 | 61.6083 | 60.3625 | 60.2719 | 60.0488 | 61.0805 | 62.1631 | 62.6949 | 61.7193 | 60.9217 | 61.3986 | 61.519 | 62.3602 | 62.6354 | 62.3641 | 62.7921 | 63.4915 | 64.1403 | 64.8055 | 62.9505 | 63.1266 | 63.7419 | 63.9783 | 63.1236 | 63.6785 | 64.0272
-2SD | 50.9167 | 66.4472 | 60.0766 | 64.5798 | 68.4609 | 71.1632 | 73.951 | 69.5516 | 70.3215 | 71.7309 | 73.0576 | 74.5228 | 75.3474 | 74.0046 | 74.8847 | 73.7515 | 73.4961 | 73.6992 | 74.5537 | 75.2992 | 75.5693 | 74.6969 | 73.9617 | 74.4657 | 74.7691 | 75.3636 | 75.7331 | 75.4151 | 75.8503 | 76.3922 | 76.9061 | 77.3855 | 75.9278 | 76.1987 | 76.6335 | 76.7333 | 76.0122 | 76.4353 | 76.6348
+2SD | 147.0833 | 137.5528 | 132.9234 | 135.8202 | 137.2058 | 133.9796 | 133.049 | 132.4484 | 130.2785 | 128.6327 | 129.609 | 129.1695 | 127.9383 | 127.1954 | 127.9903 | 127.3074 | 126.3928 | 128.3008 | 128.4463 | 127.8436 | 127.0671 | 126.6074 | 126.1217 | 126.7343 | 127.7694 | 127.3772 | 128.124 | 127.6194 | 128.083 | 127.9949 | 127.9689 | 127.7054 | 127.8369 | 128.487 | 128.1998 | 127.7532 | 127.5667 | 127.4622 | 127.0652
+3SD | 171.1249 | 155.3292 | 151.1352 | 153.6303 | 154.392 | 149.6838 | 147.8235 | 148.1726 | 145.2678 | 142.8582 | 143.7469 | 142.8311 | 141.086 | 140.4931 | 141.2667 | 140.6963 | 139.617 | 141.9512 | 141.9195 | 140.9797 | 139.9415 | 139.5851 | 139.1617 | 139.8014 | 141.0195 | 140.3806 | 141.2217 | 140.6704 | 141.1412 | 140.8956 | 140.7347 | 140.2854 | 140.8142 | 141.5591 | 141.0914 | 140.5081 | 140.4553 | 140.2189 | 139.6728

## Slide 30: Control Chart Interpretation Outright Failures
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 66 | 64 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 154 | 105 | 96 | 85 | 108 | 98
Mean | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6 | 101.6
-3SD | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259 | 51.259
-2SD | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393 | 68.0393
+2SD | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607 | 135.1607
+3SD | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941 | 151.941
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Notes: In this control chart there are a couple of instances of failure * - 2 consecutive points outwith 2 standard deviations – one on its own could be a result of random chance, remember this should occur for 1 in roughly 20 points, but the chances are much much slimmer to get 2 consecutive points outwith 2 standard deviations. This would indicate some sort of issue with the process or method that would require investigating.
1 point outwith 3 standard deviations – the likelihood of this happening by random chance is 1 in 370, very rare. It’s much more likely that there has been an issue with the process or method that would require investigating.
In either of these cases, we have to assume that samples associated with these QC results would also be affected by the same issue that caused the QC failures. We wouldn’t usually be able to report these results.

## Slide 31: Control Chart Setup Fixed Mean vs Rolling Mean
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 74 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 116 | 99 | 102 | 96.5 | 100.2 | 102.8333 | 102.5714 | 103.5 | 100.2222 | 99.6 | 99.5455 | 100.75 | 101.3077 | 101.1429 | 100.1333 | 101 | 100.1176 | 99.5556 | 100.6316 | 101.15 | 101.2381 | 101 | 100.3478 | 99.75 | 100.32 | 101 | 101.1111 | 101.6786 | 101.2759 | 101.7333 | 101.9677 | 102.2188 | 102.3333 | 101.6765 | 102.1429 | 102.2222 | 102.0541 | 101.6053 | 101.7692 | 101.675
-3SD | 26.8751 | 48.6708 | 41.8648 | 46.7697 | 51.2746 | 55.4591 | 59.1765 | 49.3375 | 51.2637 | 53.6864 | 55.2685 | 57.3466 | 58.8659 | 57.7393 | 58.7436 | 57.7724 | 57.8564 | 57.7339 | 58.8211 | 59.9632 | 60.5808 | 59.7584 | 59.0921 | 59.6102 | 59.7786 | 60.6532 | 60.9678 | 60.7727 | 61.231 | 61.9541 | 62.626 | 63.3141 | 61.5717 | 61.7745 | 62.4091 | 62.678 | 61.8881 | 62.458 | 62.8298
-2SD | 50.9167 | 66.4472 | 60.0766 | 64.5798 | 68.4609 | 71.1632 | 73.951 | 66.2991 | 67.3758 | 68.9727 | 70.429 | 72.0003 | 72.9582 | 71.8706 | 72.8291 | 71.8875 | 71.7561 | 72.0331 | 72.9307 | 73.7215 | 74.0538 | 73.2882 | 72.6447 | 73.1801 | 73.5191 | 74.1391 | 74.538 | 74.2738 | 74.7318 | 75.292 | 75.8236 | 76.3205 | 74.94 | 75.2306 | 75.6801 | 75.8034 | 75.1272 | 75.5617 | 75.7782
+2SD | 147.0833 | 137.5528 | 132.9234 | 135.8202 | 137.2058 | 133.9796 | 133.049 | 134.1453 | 131.8242 | 130.1182 | 131.071 | 130.6151 | 129.3275 | 128.396 | 129.1709 | 128.3478 | 127.355 | 129.23 | 129.3693 | 128.7547 | 127.9462 | 127.4074 | 126.8553 | 127.4599 | 128.4809 | 128.0831 | 128.8191 | 128.278 | 128.7349 | 128.6435 | 128.6139 | 128.3462 | 128.413 | 129.0551 | 128.7643 | 128.3048 | 128.0834 | 127.9767 | 127.5718
+3SD | 171.1249 | 155.3292 | 151.1352 | 153.6303 | 154.392 | 149.6838 | 147.8235 | 151.1069 | 147.9363 | 145.4045 | 146.2315 | 145.2688 | 143.4198 | 142.5274 | 143.2564 | 142.4629 | 141.2548 | 143.5293 | 143.4789 | 142.513 | 141.4192 | 140.9372 | 140.4079 | 141.0298 | 142.2214 | 141.5691 | 142.3894 | 141.779 | 142.2357 | 141.9814 | 141.8115 | 141.3526 | 141.7812 | 142.5112 | 142.0354 | 141.4301 | 141.3224 | 141.0805 | 140.5202
Notes: Replace for straight mean version

## Slide 32: Control Chart Setup Fixed Mean vs Rolling Mean
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 74 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 90 | 120 | 111 | 103 | 96 | 86 | 86 | 114 | 118 | 104 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 116 | 99 | 102 | 96.5 | 100.2 | 102.8333 | 102.5714 | 103.5 | 100.2222 | 99.6 | 99.5455 | 100.75 | 101.3077 | 101.1429 | 100.1333 | 101 | 100.1176 | 99.5556 | 100.6316 | 101.15 | 101.2381 | 101 | 100.3478 | 99.75 | 100.32 | 101 | 101.1111 | 101.6786 | 101.2759 | 101.7333 | 101.9677 | 102.2188 | 102.3333 | 101.6765 | 102.1429 | 102.2222 | 102.0541 | 101.6053 | 101.7692 | 101.675
-3SD | 26.8751 | 48.6708 | 41.8648 | 46.7697 | 51.2746 | 55.4591 | 59.1765 | 49.3375 | 51.2637 | 53.6864 | 55.2685 | 57.3466 | 58.8659 | 57.7393 | 58.7436 | 57.7724 | 57.8564 | 57.7339 | 58.8211 | 59.9632 | 60.5808 | 59.7584 | 59.0921 | 59.6102 | 59.7786 | 60.6532 | 60.9678 | 60.7727 | 61.231 | 61.9541 | 62.626 | 63.3141 | 61.5717 | 61.7745 | 62.4091 | 62.678 | 61.8881 | 62.458 | 62.8298
-2SD | 50.9167 | 66.4472 | 60.0766 | 64.5798 | 68.4609 | 71.1632 | 73.951 | 66.2991 | 67.3758 | 68.9727 | 70.429 | 72.0003 | 72.9582 | 71.8706 | 72.8291 | 71.8875 | 71.7561 | 72.0331 | 72.9307 | 73.7215 | 74.0538 | 73.2882 | 72.6447 | 73.1801 | 73.5191 | 74.1391 | 74.538 | 74.2738 | 74.7318 | 75.292 | 75.8236 | 76.3205 | 74.94 | 75.2306 | 75.6801 | 75.8034 | 75.1272 | 75.5617 | 75.7782
+2SD | 147.0833 | 137.5528 | 132.9234 | 135.8202 | 137.2058 | 133.9796 | 133.049 | 134.1453 | 131.8242 | 130.1182 | 131.071 | 130.6151 | 129.3275 | 128.396 | 129.1709 | 128.3478 | 127.355 | 129.23 | 129.3693 | 128.7547 | 127.9462 | 127.4074 | 126.8553 | 127.4599 | 128.4809 | 128.0831 | 128.8191 | 128.278 | 128.7349 | 128.6435 | 128.6139 | 128.3462 | 128.413 | 129.0551 | 128.7643 | 128.3048 | 128.0834 | 127.9767 | 127.5718
+3SD | 171.1249 | 155.3292 | 151.1352 | 153.6303 | 154.392 | 149.6838 | 147.8235 | 151.1069 | 147.9363 | 145.4045 | 146.2315 | 145.2688 | 143.4198 | 142.5274 | 143.2564 | 142.4629 | 141.2548 | 143.5293 | 143.4789 | 142.513 | 141.4192 | 140.9372 | 140.4079 | 141.0298 | 142.2214 | 141.5691 | 142.3894 | 141.779 | 142.2357 | 141.9814 | 141.8115 | 141.3526 | 141.7812 | 142.5112 | 142.0354 | 141.4301 | 141.3224 | 141.0805 | 140.5202
-2SD at end | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782 | 75.7782

## Slide 33: Control Chart Interpretation Bias
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116 | 82 | 108 | 80 | 115 | 116 | 101 | 110 | 81 | 94 | 99 | 114 | 108 | 99 | 86 | 114 | 86 | 100 | 120 | 111 | 108 | 106 | 109 | 105 | 114 | 118 | 107 | 117 | 90 | 115 | 109 | 110 | 106 | 80 | 118 | 105 | 96 | 85 | 108 | 98
Mean | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6 | 103.6
-3SD | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457 | 67.8457
-2SD | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638 | 79.7638
+2SD | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362 | 127.4362
+3SD | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543 | 139.3543
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Notes: Here we don’t have any outright failures, but I’ll draw your attention to this section of points in the middle * – here we have 10 points that are all on the same side of the mean. That’s what’s known as a bias. This situation would indicate something has gone slightly awry with the process or method, as if it was performing as expected all results should have an equal chance of falling on either side of the mean. This would need to be investigated, although rejecting associated sample results is not necessary in most cases.

## Slide 34: Control Chart Interpretation Drift
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 84 | 89 | 84 | 86 | 83 | 84 | 87 | 88 | 90 | 82 | 92 | 98 | 96 | 95 | 92 | 100 | 98 | 98 | 95 | 92 | 100 | 100 | 100 | 107 | 104 | 104 | 102 | 105 | 100 | 104 | 118 | 114 | 113 | 119 | 115 | 112 | 111 | 119 | 115 | 111
Mean | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65 | 99.65
-3SD | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449 | 66.5449
-2SD | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799 | 77.5799
+2SD | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201 | 121.7201
+3SD | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551 | 132.7551
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Notes: This is something that can be thought of as a more extreme form of bias, I hope you don’t need my * big red arrow to see the upward trend, or Drift, in these results. Again no single point has outright failed, this means that the control chart needs closer monitoring to pick up on effects like bias and drift. Especially as in real life the effect may be more subtle than I have shown here. This situation would also be cause for an investigation. It’s always good practice to rule out instrument issues but just remember for drift in particular, it is quite commonly caused by deterioration of the QC sample itself.

## Slide 35: Control Chart Interpretation Periodic Step Changes
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 116.02 | 82.16 | 108.54 | 81.28 | 117.5 | 120.32 | 107.86 | 120.24 | 95.58 | 114 | 125.62 | 148.56 | 151.94 | 99.02 | 86.16 | 114.54 | 87.28 | 92.5 | 124.32 | 117.86 | 113.24 | 110.58 | 106 | 120 | 148.56 | 161.94 | 104.02 | 117.16 | 90.54 | 116.28 | 111.5 | 114.32 | 112.86 | 90.24 | 120 | 125 | 122.62 | 135 | 151.94 | 98
Mean | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275 | 114.5275
-3SD | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846 | 55.9846
-2SD | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989 | 75.4989
+2SD | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561 | 153.5561
+3SD | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704 | 173.0704
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Notes: Here again, every point is within acceptable limits on its own (remember a single point outwith 2SD could be down to random chance) but you will notice at regular intervals in this chart*, results climb quite high then drop down suddenly. This effect can be a reflection of the routine of the method, maybe it is the preparation of new calibration standards or some instrument maintenance. This is why it is good practice to note in the QC chart when such events have occurred, so that when looking back you can match up the step change to the method event and assign a likely cause. I’m afraid to say if these events are causing a pattern in results like this, you are not preparing new calibration standards or performing the instrument maintenance *quite* often enough. Again for the purposes of demonstration, in this chart the periodic changes are quite obvious as the time interval is quite small, in your real life chart you may have to inspect more carefully to spot subtle periodic patterns.
Monitoring sugars electrode deterioration

## Slide 36: Control Chart Interpretation Analyst Variation
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
Results | 87 | 87 | 90 | 81 | 105 | 107 | 100 | 100 | 100 | 85 | 87 | 83 | 98 | 96 | 99 | 105 | 101 | 105 | 103 | 101 | 108 | 96 | 102 | 101 | 81 | 82 | 82 | 86 | 105 | 106 | 104 | 104 | 99 | 108 | 102 | 103 | 101 | 107 | 107 | 92
Mean | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4 | 97.4
-3SD | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251 | 71.1251
-2SD | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834 | 79.8834
+2SD | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166 | 114.9166
+3SD | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749 | 123.6749
Analyst B
Analyst B
Analyst B
Analyst B
Analyst C
Analyst C
Analyst C
Analyst A
Analyst A
Analyst A
μ + 3σ
μ + 2σ
μ
μ - 2σ
μ - 3σ
Notes: One final point to watch out for, this QC chart might look all in order at first glance but on further inspection * we find a pattern related to the analysts performing the method. In this instance it would be worth gathering all analysts together and finding out what Analyst A is doing differently to Analysts B and C. It might be that Analyst A requires some re-training, but they might also have unconsciously hit upon an improvement to the method that everyone else can learn from.

## Slide 37: Control Chart Management Removal of Points from Mean and SD Calculations

## Slide 38: Control Chart Management Method Uncertainty
Notes: Importance, how to calculate, use as basis for ongoing performance monitoring (reviewed annually)
NORDtest (needs ERMs or good proficiency schemes)
JC

## Slide 39: Control Chart Management Recording of issues
Notes: Root cause analysis (TAs)
Examples of issues we’ve discovered eg apparent strength seasonal bias

## Slide 40: Proficiency Schemes

## Slide 41: Standardization – z-scores
Z* scores
J scores?
Notes: JF – assigned values can depend on scheme or method. Mean/median, exclusion outliers. Simulants – (target values)
Methods – esp. MAPS NIR vs wet chem, all compared together.
Wide BIPEA ranges.

## Slide 42: Internal Round Robins
Notes: Multiple labs/sites in one company.

## Slide 43: Training
Notes: Records/signoffs.
Checking (transcription)

## Slide 44: Thank you for your attention
Hannah Ford
Analytical Research Scientist
hannah.ford@swri.co.uk
Shona Harrison
Analytical Services & Laboratory Manager
shona.harrison@swri.co.uk
Notes: And that brings us to the end, thanks very much for paying attention to me. We’ll take any questions at this point, or if you think of one later or if you’re watching the recording of this, feel free to contact me at the email above.